## Final Report
Final Report
Job 2014-0164
|  | Isotope(s) requested: | δ13C, δ15N
|  | Instrument Used: | Costech 4010 Elemental Analyzer coupled to a Thermo Delta Plus XP IRMS
|  | Analytical Code: | 010 (combustion/reduction organics)
|  | Units: | δ13C and δ15N values are reported w.r.t. VPDB and AIR respectively in parts per thousand (per mil)
|  | Principal Investigator: | Jason Stockwell
|  | Job submission contact: | Peter Eucklide
|  | Sample Material(s): | Plant and Shrimp
|  | Date Submited: | 2014-09-12
|  | Number of unknown samples analyzed: | 7
|  | Number of reference samples analyzed: | 8
|  | Quality Control Reference Material 1: | 36-UWSIF-Glutamic 1
|  | Quality Control Reference Material 2: | 39-UWSIF-Glutamic 2
|  | Quality Assessment Reference Material 3: | UWSIF05 (alfalfa)
Quality Control Data |  |  |  |  | Quality Assesment Data
Reference Material 1 | Known |  | Known |  | Reference Material 3 | Known | Normalized | Known | Normalized
36-UWSIF-Glutamic 1 | δ13C VPDB |  | δ15N AIR |  | UWSIF05 (alfalfa) | δ13C VPDB | δ13C VPDB | δ15N AIR | δ15N AIR
average (N=2) | -28.28 |  | -4.62 |  | average (N=4) | -27.52 | -27.5684 | 0.21 | 0.10104
standard uncertainty | 0.027329 |  | 0.039004 |  | standard uncertainty |  | 0.044389 |  | 0.038842
Reference Material 2 | Known |  | Known |  | Long-Term | δ13C VPDB | -27.82 | δ15N AIR | -0.19
39-UWSIF-Glutamic 2 | δ13C VPDB |  | δ15N AIR |  | Acceptable Range | 2 σ = 0.3 | -27.22 | 2 σ = 0.4 | 0.61
average (N=2) | 24.36 |  | 27.89 |  | Weight Percent | Known | Measured | Known | Measured
standard uncertainty | 0.017519 |  | 0.021984 |  |  | Wt% C | Wt% C | Wt% N | Wt% N
[5 empty] | average (N=4) | 42.99 | 41.0302 | 3.25 | 3.2488
[5 empty] | standard uncertainty |  | 0.085827 |  | 0.022336
[5 empty] | relative error (%) |  | 4.5587 |  | 0.48391
Record Keeping |  |  |  |  | Quality Assurance Approval
Date Reported: | 2014-10-07 |  |  |  | Reviewer: | Craig Cook
Date Invoiced: | 2014-10-07 |  |  |  | Title: | Facility Director
Initial: | cjm |  |  |  | Date Reviewed: | 2014-10-07
Comments: |  |  |  |  | Comments: | Data meet all QAQC criteria in the SOP.
Analytical Comments:
*Sample weight percents are calculated using the sample weights reported by the user and are dependent upon the accuracy of these values.
Quality Control Color Legend
Yellow | Peak amplitudes too low for reliable results. Use with extreme caution or rerun the sample.
Green | Weight percent values wrong. Use with caution.
Pink | Potential outlier. Use with caution.
Blue | Sample lost during the analysis. Reload and rerun the sample.
Olive | Possible sample ID problem. Check loading documents.
SIF ID | Sample ID | δ15N | δ13C | Wt% N* | Wt% C* | C:N ratio | Comments
20140164.001 | m129 | -10.4786 | -31.1038 | 2.2009 | 51.3923 | 23.3509
20140164.002 | m129b | -10.4073 | -31.0592 | 2.1775 | 50.1603 | 23.0352
20140164.003 | m129c | -10.5066 | -31.0929 | 2.149 | 50.5359 | 23.5157
20140164.004 | m93d | -8.224 | -21.8548 | 2.4144 | 47.2441 | 19.5679
20140164.005 | m93e | -8.3163 | -21.7953 | 2.3869 | 46.5805 | 19.5155
20140164.006 | m99d | 13.3283 | -21.7646 | 9.5393 | 41.63 | 4.364
20140164.007 | m99e | 13.2641 | -21.726 | 8.7925 | 37.9719 | 4.3187
Stable Isotope Facility

## QAQC, calculations
Line | Identifier 1 | δ15N | δ13C | Line | Identifier 1 | δ15N | δ13C | Line | Identifier 1 | Absolute Difference |  | Weight % measured |  | δ15N |  | δ13C
|  | Normalized | Normalized |  |  | Normalized | Normalized |  |  | % Nitrogen | % Carbon | % Nitrogen | % Carbon | normalized |  | normalized
QC Reference Materials |  |  |  | QC Reference Materials |  |  |  | QS Reference Material
4 | 36-UWSIF-UT Glut 120140164.14 | -4.6316 | -28.2913 | 6 | 39-UWSIF-UW Glut 220140164.11 | 27.8613 | 24.3813 | 8 | 05-UWSIF-20140164.11 | 0.0052548 | 1.9497 | 3.2553 | 41.0403 |  | 0.15419 | -27.5114
5 | 36-UWSIF-UT Glut 120140164.15 | -4.6006 | -28.2666 | 7 | 39-UWSIF-UW Glut 220140164.12 | 27.9165 | 24.3426 | 9 | 05-UWSIF-20140164.12 | 0.016831 | 2.0299 | 3.2668 | 40.9601 |  | 0.10004 | -27.6125
2 | average | -4.6161 | -28.279 | 2 | average | 27.8889 | 24.362 | 17 | 05-UWSIF-0.13 | 0.0070317 | 2.0171 | 3.257 | 40.9729 |  | 0.088001 | -27.557
| std uncertainty | 0.021984 | 0.017519 |  | std uncertainty | 0.039004 | 0.027329 | 18 | 05-UWSIF-0.14 | 0.033791 | 1.8423 | 3.2162 | 41.1477 |  | 0.061925 | -27.5927
| δ15N known | -4.62 |  |  | δ15N known | 27.89 |  | 4 |  |  | Known % | 3.25 | 42.99 | δ15N known | 0.21 | δ13C known | -27.52
| δ13C known |  | -28.28 |  | δ13C known |  | 24.36 |  |  |  | Average | 3.2488 | 41.0302 | average δ15N | 0.10104 | average δ13C | -27.5684
[11 empty] | Stdev | 0.022336 | 0.085827 | stdev δ15N | 0.038842 | stdev δ13C | 0.044389
[11 empty] | Absolute Error | 0.015727 | 1.9598 | Long-Term | 0.61 | Long-Term | -27.22
[11 empty] | Relative Error | 0.48391 | 4.5587 | Acceptable Range | -0.19 | Acceptable Range | -27.82

## Run 1
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | 2.4 | 1933 | -9.915 | 7736 | 3.978 | 2.2009 | 51.3923 | 23.3509 | -10.4786 | -31.1038
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | 2.7 | 2179 | -9.844 | 8466 | 4.023 | 2.1775 | 50.1603 | 23.0352 | -10.4073 | -31.0592
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | 2.4 | 1888 | -9.943 | 7606 | 3.989 | 2.149 | 50.5359 | 23.5157 | -10.5066 | -31.0929
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d | 2.4 | 2084 | -7.667 | 6995 | 13.311 | 2.4144 | 47.2441 | 19.5679 | -8.224 | -21.8548
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e | 2.5 | 2161 | -7.759 | 7204 | 13.371 | 2.3869 | 46.5805 | 19.5155 | -8.3163 | -21.7953
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d | 1.1 | 3563 | 13.823 | 2699 | 13.402 | 9.5393 | 41.63 | 4.364 | 13.3283 | -21.7646
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e | 1.2 | 3589 | 13.759 | 2689 | 13.441 | 8.7925 | 37.9719 | 4.3187 | 13.2641 | -21.726
Reference Check
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | 2.149 | 2420 | 0.687 | 5397 | 7.603 | 3.2553 | 41.0403 | 12.6074 | 0.15419 | -27.5114
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | 2.186 | 2478 | 0.633 | 5489 | 7.501 | 3.2668 | 40.9601 | 12.5382 | 0.10004 | -27.6125
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- | 2.226 | 2522 | 0.621 | 5598 | 7.557 | 3.257 | 40.9729 | 12.5798 | 0.088001 | -27.557
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- | 2.178 | 2446 | 0.595 | 5526 | 7.521 | 3.2162 | 41.1477 | 12.7938 | 0.061925 | -27.5927
[5 empty] | average | 0.634 |  | 7.5455 | 3.2488 | 41.0302 |  | 0.10104 | -27.5684
[5 empty] | std. dev. | 0.03873 |  | 0.044792 | 0.022336 | 0.085827 |  | 0.038842 | 0.044389
Reference Material
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | 0.827 | 2586 | -4.085 | 1948 | 6.816 | 9.52 | 40.81 | 4.2868 | -4.6316 | -28.2913
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | 1.022 | 3286 | -4.054 | 2504 | 6.841 | 9.52 | 40.81 | 4.2868 | -4.6006 | -28.2666
[5 empty] | average | -4.0695 |  | 6.8285 | 9.52 | 40.81 |  | -4.6161 | -28.279
[5 empty] | std. dev. | 0.02192 |  | 0.017678 | 0 | 0 |  | 0.021984 | 0.017519
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | 0.783 | 2677 | 28.314 | 2026 | 59.967 | 10.3208 | 44.2661 | 4.289 | 27.8613 | 24.3813
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | 0.814 | 2778 | 28.369 | 2077 | 59.928 | 10.3632 | 44.3839 | 4.2828 | 27.9165 | 24.3426
[5 empty] | average | 28.3415 |  | 59.9475 | 10.342 | 44.325 |  | 27.8889 | 24.362
[5 empty] | std. dev. | 0.038891 |  | 0.027577 | 0.029979 | 0.083262 |  | 0.039004 | 0.027329
15N Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -4.0695 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 28.3415 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual
|  | UWSIF05 (Alfalfa) | 0.10104 | 0.21 | 0.10896
Percentage
|  | UWSIF05 (Alfalfa) | 3.2488 | 3.255
13C Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | 6.8285 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 59.9475 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual
|  | UWSIF05 (Alfalfa) | -27.5684 | -27.52 | 0.04841
Percentage
|  | UWSIF05 (Alfalfa) | 41.0302 | 42.995
Date Analyzed:9/30/2014
Analyst:DEW
[6 empty] | Instrument: | Costech EA 4010

## Original 1
Time Code | Row | Identifier 1 | Identifier 2 | Comment | Amount | Peak Nr | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Amt% | Area All | Area 28 | BGD 28 | BGD 29 | BGD 30 | Area 44 | BGD 44 | BGD 45 | BGD 46 | AT% 13C/12C | AT% 15N/14N | rR 45CO2/44CO2 | rR 29N2/28N2
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 1 | 3522 | -0.012 |  |  | 11.2706 | 69.519 | 68.982 | 12.8 | 9.5 | 326.0 | [5 empty] | 0.36647 |  | 0.7284
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 2 | 3520 | 0 |  |  | 11.2624 | 69.47 | 68.933 | 12.7 | 9.5 | 333.7 | [5 empty] | 0.36647 |  | 0.72841
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 3 | 2372 | -4.016 |  |  | 9.5983 | 59.261 | 58.803 | 12.7 | 9.4 | 311.3 | [5 empty] | 0.36501 |  | 0.72549
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 4 |  |  | 1836 | 6.856 | 43.3089 | 50.283 |  |  |  |  | 49.5 | 0.8 | 0.9 | 3.4 | 1.1132 |  | 1.1581
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 5 |  |  | 4102 | 0.176 | 68.6374 | 81.062 |  |  |  |  | 79.811 | 2.7 | 3.1 | 6.2 | 1.1059 |  | 1.1502
2014/09/30 13:15:33 | 1 | 36-UWSIF-UT Glut 120140164.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0164 Run 1 Stockwell | 0.776 | 6 |  |  | 4100 | 0 | 68.7699 | 81.226 |  |  |  |  | 79.973 | 4.6 | 5.4 | 8.9 | 1.1057 |  | 1.15
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 1 | 3523 | 0.008 |  |  | 11.1274 | 69.521 | 68.984 | 13.0 | 9.8 | 318.1 | [5 empty] | 0.36647 |  | 0.7285
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 2 | 3523 | 0 |  |  | 11.1228 | 69.492 | 68.956 | 12.9 | 9.6 | 327.6 | [5 empty] | 0.36647 |  | 0.7285
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 3 | 2444 | -3.964 |  |  | 9.5733 | 59.865 | 59.402 | 12.9 | 9.6 | 307.5 | [5 empty] | 0.36502 |  | 0.72561
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 4 |  |  | 1889 | 6.83 | 43.144 | 50.75 |  |  |  |  | 49.959 | 1.0 | 1.1 | 3.7 | 1.1131 |  | 1.159
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 5 |  |  | 4099 | 0.172 | 67.7767 | 81.077 |  |  |  |  | 79.826 | 2.7 | 3.2 | 6.2 | 1.1058 |  | 1.1511
2014/09/30 13:24:45 | 2 | 36-UWSIF-UT Glut 120140164.12 | 36-UWSIF-UT Glut 1 | Analyst:DEW | 0.786 | 6 |  |  | 4100 | 0 | 67.8616 | 81.184 |  |  |  |  | 79.931 | 4.6 | 5.4 | 8.9 | 1.1057 |  | 1.1509
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 1 | 3524 | 0.037 |  |  |  | 69.478 | 68.941 | 13.1 | 9.9 | 317.3 | [5 empty] | 0.36649 |  | 0.72854
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 2 | 3522 | 0 |  |  |  | 69.461 | 68.924 | 13.1 | 9.8 | 327.3 | [5 empty] | 0.36647 |  | 0.72851
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 3 | 1312 | -4.062 |  |  | 9.52 | 33.305 | 33.05 | 13.1 | 9.7 | 306.6 | [5 empty] | 0.36499 |  | 0.72555
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 4 |  |  | 987 | 6.911 | 40.81 | 27.673 |  |  |  |  | 27.241 | 1.0 | 1.2 | 3.8 | 1.1132 |  | 1.1594
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 5 |  |  | 4100 | 0.153 |  | 81.045 |  |  |  |  | 79.794 | 2.4 | 2.8 | 5.7 | 1.1058 |  | 1.1514
2014/09/30 13:33:58 | 3 | 36-UWSIF-UT Glut 120140164.13 | 36-UWSIF-UT Glut 1 | Samples on Column: 024 | 0.434 | 6 |  |  | 4102 | 0 |  | 81.255 |  |  |  |  | 80.001 | 4.5 | 5.2 | 8.7 | 1.1057 |  | 1.1512
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 1 | 3524 | -0.003 |  |  |  | 69.542 | 69.005 | 13.1 | 9.9 | 315.1 | [5 empty] | 0.36647 |  | 0.72854
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 2 | 3523 | 0 |  |  |  | 69.453 | 68.917 | 13.0 | 9.8 | 325.6 | [5 empty] | 0.36647 |  | 0.72854
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 3 | 2586 | -4.085 |  |  | 9.52 | 63.239 | 62.75 | 13.1 | 9.8 | 305.7 | [5 empty] | 0.36498 |  | 0.72556
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 4 |  |  | 1948 | 6.816 | 40.81 | 53.48 |  |  |  |  | 52.645 | 1.0 | 1.2 | 3.7 | 1.1131 |  | 1.1595
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 5 |  |  | 4100 | 0.19 |  | 81.026 |  |  |  |  | 79.775 | 3.0 | 3.5 | 6.5 | 1.1059 |  | 1.1516
2014/09/30 13:43:10 | 4 | 36-UWSIF-UT Glut 120140164.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 400 | 0.827 | 6 |  |  | 4101 | 0 |  | 81.184 |  |  |  |  | 79.931 | 4.7 | 5.5 | 9.0 | 1.1057 |  | 1.1513
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 1 | 3520 | 0.013 |  |  |  | 69.464 | 68.928 | 13.3 | 10.0 | 318.0 | [5 empty] | 0.36648 |  | 0.72854
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 2 | 3521 | 0 |  |  |  | 69.394 | 68.858 | 13.1 | 9.9 | 327.9 | [5 empty] | 0.36647 |  | 0.72853
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 3 | 3286 | -4.054 |  |  | 9.52 | 78.264 | 77.657 | 13.2 | 9.9 | 308.6 | [5 empty] | 0.36499 |  | 0.72558
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 4 |  |  | 2504 | 6.841 | 40.81 | 66.694 |  |  |  |  | 65.654 | 1.1 | 1.2 | 3.9 | 1.1131 |  | 1.1595
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 5 |  |  | 4101 | 0.141 |  | 81.086 |  |  |  |  | 79.834 | 3.2 | 3.8 | 7.0 | 1.1058 |  | 1.1515
2014/09/30 13:52:22 | 5 | 36-UWSIF-UT Glut 120140164.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 1400 | 1.022 | 6 |  |  | 4100 | 0 |  | 81.168 |  |  |  |  | 79.915 | 4.9 | 5.6 | 9.3 | 1.1057 |  | 1.1513
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 1 | 3522 | 0.004 |  |  | 11.0491 | 69.595 | 69.057 | 13.3 | 10.1 | 320.2 | [5 empty] | 0.36647 |  | 0.72851
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 2 | 3523 | 0 |  |  | 11.0321 | 69.489 | 68.952 | 13.2 | 9.9 | 330.4 | [5 empty] | 0.36647 |  | 0.72851
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 3 | 2677 | 28.314 |  |  | 10.3208 | 65.024 | 64.506 | 13.3 | 10.0 | 310.9 | [5 empty] | 0.37681 |  | 0.74914
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 4 |  |  | 2026 | 59.967 | 44.2661 | 55.039 |  |  |  |  | 54.15 | 1.1 | 1.3 | 4.0 | 1.1712 |  | 1.2168
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 5 |  |  | 4098 | 0.329 | 64.1781 | 81.057 |  |  |  |  | 79.805 | 3.1 | 3.7 | 6.7 | 1.106 |  | 1.1519
2014/09/30 14:01:40 | 6 | 39-UWSIF-UW Glut 220140164.11 | 39-UWSIF-UW Glut 2 | Mass 28:11 | 0.783 | 6 |  |  | 4097 | 0 | 64.2424 | 81.143 |  |  |  |  | 79.89 | 4.8 | 5.6 | 9.1 | 1.1057 |  | 1.1515
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 1 | 3522 | 0.038 |  |  | 10.6134 | 69.498 | 68.961 | 13.3 | 10.1 | 320.6 | [5 empty] | 0.36649 |  | 0.72854
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 2 | 3519 | 0 |  |  | 10.6016 | 69.421 | 68.885 | 13.2 | 10.0 | 330.7 | [5 empty] | 0.36647 |  | 0.72852
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 3 | 2778 | 28.369 |  |  | 10.3632 | 67.866 | 67.324 | 13.3 | 10.0 | 310.8 | [5 empty] | 0.37683 |  | 0.74918
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 4 |  |  | 2077 | 59.928 | 44.3839 | 57.454 |  |  |  |  | 56.525 | 1.1 | 1.4 | 4.0 | 1.1711 |  | 1.2167
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 5 |  |  | 4098 | 0.331 | 61.7693 | 81.106 |  |  |  |  | 79.854 | 3.2 | 3.8 | 6.9 | 1.106 |  | 1.1518
2014/09/30 14:10:59 | 7 | 39-UWSIF-UW Glut 220140164.12 | 39-UWSIF-UW Glut 2 | Mass 29:8 | 0.814 | 6 |  |  | 4097 | 0 | 61.7955 | 81.142 |  |  |  |  | 79.89 | 4.8 | 5.7 | 9.2 | 1.1057 |  | 1.1515
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 1 | 3524 | 0.012 |  |  | 4.0229 | 69.545 | 69.008 | 13.4 | 10.2 | 322.3 | [5 empty] | 0.36648 |  | 0.72842
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 2 | 3523 | 0 |  |  | 4.0219 | 69.528 | 68.991 | 13.3 | 10.0 | 332.5 | [5 empty] | 0.36647 |  | 0.72841
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 3 | 2420 | 0.687 |  |  | 3.2553 | 56.317 | 55.88 | 13.4 | 10.0 | 312.1 | [5 empty] | 0.36672 |  | 0.72891
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 4 |  |  | 5397 | 7.603 | 41.0403 | 148.316 |  |  |  |  | 146 | 1.1 | 1.4 | 4.0 | 1.114 |  | 1.1605
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 5 |  |  | 4098 | 0.086 | 23.3716 | 81.013 |  |  |  |  | 79.763 | 4.9 | 5.7 | 9.3 | 1.1058 |  | 1.1516
2014/09/30 14:20:12 | 8 | 05-UWSIF-20140164.11 | 05-UWSIF- | Mass 30:453 | 2.149 | 6 |  |  | 4098 | 0 | 23.4047 | 81.134 |  |  |  |  | 79.882 | 5.6 | 6.5 | 10.4 | 1.1057 |  | 1.1514
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 1 | 3524 | 0.02 |  |  | 3.9546 | 69.542 | 69.005 | 13.7 | 10.5 | 324.3 | [5 empty] | 0.36648 |  | 0.72842
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 2 | 3521 | 0 |  |  | 3.9525 | 69.504 | 68.967 | 13.5 | 10.3 | 334.3 | [5 empty] | 0.36647 |  | 0.72841
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 3 | 2478 | 0.633 |  |  | 3.2668 | 57.486 | 57.04 | 13.6 | 10.3 | 313.9 | [5 empty] | 0.3667 |  | 0.72887
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 4 |  |  | 5489 | 7.501 | 40.9601 | 150.813 |  |  |  |  | 148.458 | 1.4 | 1.7 | 4.3 | 1.1139 |  | 1.1603
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 5 |  |  | 4095 | 0.114 | 22.9607 | 80.957 |  |  |  |  | 79.707 | 5.0 | 5.8 | 9.5 | 1.1058 |  | 1.1515
2014/09/30 14:29:24 | 9 | 05-UWSIF-20140164.12 | 05-UWSIF- | Peak Center:3.064 | 2.186 | 6 |  |  | 4094 | 0 | 23.0033 | 81.115 |  |  |  |  | 79.863 | 5.7 | 6.6 | 10.5 | 1.1057 |  | 1.1514
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 1 | 3522 | 0.035 |  |  | 3.6012 | 69.527 | 68.99 | 13.8 | 10.5 | 325.8 | [5 empty] | 0.36649 |  | 0.72844
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 2 | 3519 | 0 |  |  | 3.5976 | 69.457 | 68.92 | 13.6 | 10.3 | 335.6 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 3 | 1933 | -9.915 |  |  | 2.2009 | 42.556 | 42.23 | 13.7 | 10.4 | 315.0 | [5 empty] | 0.36285 |  | 0.72119
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 4 |  |  | 7736 | 3.978 | 51.3923 | 216.815 |  |  |  |  | 213.438 | 1.5 | 1.7 | 4.4 | 1.11 |  | 1.1565
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 5 |  |  | 4094 | 0.054 | 20.929 | 81.02 |  |  |  |  | 79.769 | 6.1 | 7.1 | 11.2 | 1.1057 |  | 1.1514
2014/09/30 14:38:37 | 10 | 20140164.001 | m129 | Instrument: ANNIE | 2.4 | 6 |  |  | 4093 | 0 | 20.9408 | 81.068 |  |  |  |  | 79.817 | 6.3 | 7.3 | 11.3 | 1.1057 |  | 1.1514
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 1 | 3521 | -0.002 |  |  | 3.1983 | 69.466 | 68.93 | 13.9 | 10.8 | 326.1 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 2 | 3521 | 0 |  |  | 3.1973 | 69.445 | 68.908 | 13.9 | 10.6 | 336.6 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 3 | 2179 | -9.844 |  |  | 2.1775 | 47.356 | 46.992 | 13.9 | 10.5 | 315.3 | [5 empty] | 0.36288 |  | 0.72125
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 4 |  |  | 8466 | 4.023 | 50.1603 | 242.122 |  |  |  |  | 238.351 | 1.6 | 1.9 | 4.7 | 1.1101 |  | 1.1564
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 5 |  |  | 4098 | 0.043 | 18.6116 | 81.057 |  |  |  |  | 79.806 | 6.7 | 7.8 | 12.0 | 1.1057 |  | 1.1512
2014/09/30 14:47:49 | 11 | 20140164.002 | m129b | Data analyst: | 2.7 | 6 |  |  | 4094 | 0 | 18.6216 | 81.103 |  |  |  |  | 79.851 | 6.6 | 7.6 | 11.7 | 1.1057 |  | 1.1512
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 1 | 3521 | 0.024 |  |  | 3.5996 | 69.495 | 68.958 | 14.1 | 10.9 | 327.7 | [5 empty] | 0.36648 |  | 0.72844
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 2 | 3521 | 0 |  |  | 3.5982 | 69.468 | 68.932 | 13.8 | 10.6 | 337.6 | [5 empty] | 0.36647 |  | 0.72843
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 3 | 1888 | -9.943 |  |  | 2.149 | 41.556 | 41.237 | 13.9 | 10.6 | 316.6 | [5 empty] | 0.36284 |  | 0.72118
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 4 |  |  | 7606 | 3.989 | 50.5359 | 212.611 |  |  |  |  | 209.3 | 1.7 | 2.0 | 4.9 | 1.11 |  | 1.1565
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 5 |  |  | 4100 | 0.068 | 20.9646 | 81.165 |  |  |  |  | 79.912 | 6.2 | 7.2 | 11.3 | 1.1057 |  | 1.1514
2014/09/30 14:57:02 | 12 | 20140164.003 | m129c | Comments: | 2.4 | 6 |  |  | 4097 | 0 | 20.9738 | 81.203 |  |  |  |  | 79.949 | 6.4 | 7.4 | 11.4 | 1.1057 |  | 1.1513
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 1 | 3523 | 0.023 |  |  | 3.6008 | 69.52 | 68.982 | 13.9 | 10.8 | 327.2 | [5 empty] | 0.36648 |  | 0.72843
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 2 | 3522 | 0 |  |  | 3.6008 | 69.518 | 68.981 | 13.8 | 10.6 | 337.7 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 3 | 2084 | -7.667 |  |  | 2.4144 | 46.674 | 46.314 | 13.8 | 10.5 | 317.1 | [5 empty] | 0.36367 |  | 0.72283
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 4 |  |  | 6995 | 13.311 | 47.2441 | 196.698 |  |  |  |  | 193.614 | 1.7 | 2.0 | 4.9 | 1.1202 |  | 1.1666
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 5 |  |  | 4099 | 0.09 | 20.9686 | 81.181 |  |  |  |  | 79.928 | 6.0 | 7.0 | 11.0 | 1.1058 |  | 1.1515
2014/09/30 15:06:15 | 13 | 20140164.004 | m93d |  | 2.4 | 6 |  |  | 4096 | 0 | 20.9657 | 81.17 |  |  |  |  | 79.916 | 6.3 | 7.3 | 11.3 | 1.1057 |  | 1.1514
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 1 | 3524 | -0.011 |  |  | 3.4569 | 69.521 | 68.984 | 13.5 | 10.5 | 328.0 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 2 | 3523 | 0 |  |  | 3.4577 | 69.537 | 69 | 13.5 | 10.3 | 338.4 | [5 empty] | 0.36647 |  | 0.72843
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 3 | 2161 | -7.759 |  |  | 2.3869 | 48.06 | 47.69 | 13.6 | 10.4 | 317.3 | [5 empty] | 0.36364 |  | 0.72277
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 4 |  |  | 7204 | 13.371 | 46.5805 | 202.827 |  |  |  |  | 199.647 | 1.7 | 2.0 | 4.9 | 1.1203 |  | 1.1667
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 5 |  |  | 4097 | 0.127 | 20.1006 | 81.057 |  |  |  |  | 79.806 | 6.1 | 7.2 | 11.2 | 1.1058 |  | 1.1516
2014/09/30 15:15:28 | 14 | 20140164.005 | m93e |  | 2.5 | 6 |  |  | 4097 | 0 | 20.1289 | 81.177 |  |  |  |  | 79.924 | 6.3 | 7.4 | 11.5 | 1.1057 |  | 1.1514
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 1 | 3521 | 0.028 |  |  | 7.8553 | 69.51 | 68.973 | 14.0 | 10.8 | 328.8 | [5 empty] | 0.36648 |  | 0.72844
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 2 | 3520 | 0 |  |  | 7.8501 | 69.464 | 68.927 | 13.8 | 10.6 | 339.3 | [5 empty] | 0.36647 |  | 0.72842
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 3 | 3563 | 13.823 |  |  | 9.5393 | 84.34 | 83.675 | 13.8 | 10.5 | 319.7 | [5 empty] | 0.37152 |  | 0.73849
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 4 |  |  | 2699 | 13.402 | 41.63 | 73.53 |  |  |  |  | 72.376 | 1.7 | 2.0 | 4.9 | 1.1203 |  | 1.1669
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 5 |  |  | 4096 | 0.139 | 45.6761 | 81.044 |  |  |  |  | 79.793 | 3.8 | 4.4 | 7.8 | 1.1058 |  | 1.1516
2014/09/30 15:24:40 | 15 | 20140164.006 | m99d |  | 1.1 | 6 |  |  | 4097 | 0 | 45.7066 | 81.101 |  |  |  |  | 79.849 | 5.2 | 6.1 | 9.9 | 1.1057 |  | 1.1515
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 1 | 3521 | -0.003 |  |  | 7.1956 | 69.462 | 68.925 | 13.7 | 10.5 | 330.8 | [5 empty] | 0.36647 |  | 0.72845
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 2 | 3521 | 0 |  |  | 7.1914 | 69.421 | 68.884 | 13.6 | 10.3 | 340.8 | [5 empty] | 0.36647 |  | 0.72845
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 3 | 3589 | 13.759 |  |  | 8.7925 | 84.802 | 84.133 | 13.7 | 10.3 | 320.8 | [5 empty] | 0.3715 |  | 0.73848
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 4 |  |  | 2689 | 13.441 | 37.9719 | 73.149 |  |  |  |  | 72.001 | 1.4 | 1.7 | 4.5 | 1.1204 |  | 1.1671
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 5 |  |  | 4096 | 0.137 | 41.8431 | 80.99 |  |  |  |  | 79.739 | 3.7 | 4.3 | 7.7 | 1.1058 |  | 1.1518
2014/09/30 15:33:53 | 16 | 20140164.007 | m99e |  | 1.2 | 6 |  |  | 4097 | 0 | 41.8796 | 81.064 |  |  |  |  | 79.813 | 5.1 | 6.0 | 9.8 | 1.1057 |  | 1.1516
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 1 | 3522 | -0.01 |  |  | 3.8848 | 69.563 | 69.026 | 13.6 | 10.4 | 331.6 | [5 empty] | 0.36647 |  | 0.72845
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 2 | 3520 | 0 |  |  | 3.8799 | 69.476 | 68.939 | 13.5 | 10.2 | 341.7 | [5 empty] | 0.36647 |  | 0.72846
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 3 | 2522 | 0.621 |  |  | 3.257 | 58.36 | 57.906 | 13.6 | 10.3 | 321.3 | [5 empty] | 0.3667 |  | 0.72891
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 4 |  |  | 5598 | 7.557 | 40.9729 | 153.924 |  |  |  |  | 151.518 | 1.4 | 1.6 | 4.3 | 1.1139 |  | 1.1606
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 5 |  |  | 4098 | 0.114 | 22.5755 | 81.06 |  |  |  |  | 79.808 | 5.1 | 5.9 | 9.7 | 1.1058 |  | 1.1516
2014/09/30 15:43:12 | 17 | 05-UWSIF-0.13 | 05-UWSIF- |  | 2.226 | 6 |  |  | 4096 | 0 | 22.5887 | 81.11 |  |  |  |  | 79.858 | 5.8 | 6.7 | 10.6 | 1.1057 |  | 1.1515
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 1 | 3521 | 0.054 |  |  | 3.965 | 69.469 | 68.932 | 13.8 | 10.6 | 332.7 | [5 empty] | 0.36649 |  | 0.7285
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 2 | 3520 | 0 |  |  | 3.9603 | 69.387 | 68.851 | 13.7 | 10.5 | 342.8 | [5 empty] | 0.36647 |  | 0.72846
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 3 | 2446 | 0.595 |  |  | 3.2162 | 56.392 | 55.954 | 13.7 | 10.4 | 322.4 | [5 empty] | 0.36669 |  | 0.7289
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 4 |  |  | 5526 | 7.521 | 41.1477 | 150.963 |  |  |  |  | 148.605 | 1.5 | 1.8 | 4.6 | 1.1139 |  | 1.1605
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 5 |  |  | 4097 | 0.088 | 23.0671 | 81.038 |  |  |  |  | 79.787 | 5.1 | 5.9 | 9.7 | 1.1058 |  | 1.1515
2014/09/30 15:52:25 | 18 | 05-UWSIF-0.14 | 05-UWSIF- |  | 2.178 | 6 |  |  | 4097 | 0 | 23.1072 | 81.186 |  |  |  |  | 79.933 | 5.8 | 6.8 | 10.7 | 1.1057 |  | 1.1514

## Analysis Information
Principal of operation
●The Finnigan DeltaPlus XP is run in continuous flow mode and connected to either a Carlo Erba 1110 or a Costech 4010 elemental analyzer via a Finnigan ConFlo III interface.
●A sample contained within a tin capsule is dropped into a combustion reactor held at 1020°C. For nitrogen and carbon analysis, the combustion reactor contains chromium oxide for oxidation and silvered cobaltous/cobaltic oxide for removal of sulfur. When the tin capsule is exposed to a gas flow temporarily enriched with ultra-high purity oxygen (25-30 mL min-1), flash combustion occurs which raises the temperature of the sample to >1700°C. The encapsulated sample, depending on its composition, combusts generating one or more of these gases: N2, NxOx, CO2, and H2O. The reduction reactor contains reduced copper wires for the reduction of nitrogen oxides to N2 and the removal of excess O2. An adsorption trap containing magnesium perchlorate removes the H2O. The remaining N2 and CO2 gases travel through a Porapak Q chromatographic column (80-100 mL min-1 at 50˚C) as the carrier gas and then moves to the ConFlo III open-split interface.
●The gas then moves through the ConFlo III to the Finnigan DeltaPlus XP mass spectrometer. Reference gases are pulsed by the ConFlo III into the mass spectrometer for proper mass balance.
QA/QC
Quality assesment of carbon and nitrogen isotope composition of solids is based on the standard uncertainty of the known value of the secondary laboratory reference material calculated on multiple analyses. For carbon isotope composition, if the standard uncertainty is greater than 0.15‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.3‰). The carbon isotopic composition is reported in per mil relative to VPDB scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST 8542 sucrose ANU, respectively are -26.24‰, +37.76‰, and -10.05‰. For nitrogen isotope ratio composition, if the standard uncertainty is greater than 0.2‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.4‰). The nitrogen isotopic composition is reported in per mil relative to nitrogen in air on a scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST8549 potassium nitrate, respectively are -4.52‰, +47.57‰, and 4.36‰.
Normalization
Whenever isotopic analyses are performed, reference materials must be included with the unknowns. In general two ‘bracketing’ reference materials should be used for correcting purposes. These reference materials ideally would be similar in chemical complexity to the unknowns but have very different isotopic compositions. Using two such reference materials, you can generate two linear equations, one for each referenece material. Since these linear equations will have the same slope and intercept, combining the two equations will generate a single linear equation that can be used to correct the unknowns.
The linear equation is derived as follows:
Slope (m): m=[δRM 1-known – δRM 2-known] / [δRM 1-measured – δRM 2-measured]
Intercept (b): b= δRM 1-known – [m* δRM 1-measured]
Normalization equation:
δSA-COR = m* δSA-measured + b
Where:
δRM 1-known = the accepted delta value of reference material 1
δRM 2-known = the accepted delta value of reference material 2
δRM 1-measured = the measured delta value of reference material 1
δRM 2-measured = the measured delta value of reference material 2
δSA-measured = measured delta value of sample
δSA-COR = corrected delta value of sample
13C and 15N analyses of solids

## Contact
For questions about the analysis, please contact:
|  | Chandelle Macdonald | Master Technician
| email: | cmacdon1@uwyo.edu
| Phone: | (307) 766-6373
Contact information for UW Stable Isotope Facility:
| Personnel: | Dr. David G. Williams | Faculty Director
dgw@uwyo.edu
|  | Craig Cook | Facility Director
ccook21@uwyo.edu
|  | Chandelle Macdonald | Master Technician
cmacdon1@uwyo.edu
| Address: | UWYO Stable Isotope Facility
University of Wyoming
Berry Biodiversity Center Rm 214
Laramie, WY 82071
| email: | uwyosif@uwyo.edu
| Phone: | (307) 766-6373
| Fax: | (307) 766-6403